## Projects
Analysis of Open Source COVID-19 Pandemic Ventilator Projects
Look Down! We've added tabs for modules to encourage modularity!
2020-04-15 | Public Invention | https://www.pubinv.org |  |  | Home Repo: |  | https://github.com/PubInv/covid19-vent-list
Link to definition of evaluation criteria: | https://docs.google.com/document/d/e/2PACX-1vRl9yZ27KvslftcNvweHgH1A81pO8gHL62TWpY_VY-UELWdK9x-4-3hNw3DbkemClzExPsg8RfnxilP/pub
Project Name | Project Link | Openness | Buildability (1 unit) | Community Support | Functional Testing | Reliability Testing | COVID-19 Suitability | Clinician Friendly | Average | Manufacturability (1000s) | Date Last Evaluated | Point of Contact | Team Needs | Drive | Spontaneous Breathing | Fraction O2 Control (FiO2) | Existing Device or New Design | Notes | Languages | Microcontroller | Licensing notes | Software Notes | FDA (or other gov.) Emergency Approval
Medtronic Puritan Bennett (PB) 560 | http://newsroom.medtronic.com/news-releases/news-release-details/medtronic-shares-ventilation-design-specifications-accelerate | 4 | 2.5 | 4 | 5 | 5 | 4 | 5 | 4.2143 | 4 | 2020-04-08 |  |  | Pump |  |  | Existing | Medtronic is a major manufacturer. Release appears to be detailed and complete. Open license is time-limited but good for the pandemic. May be hard to build outside of major manufacturing. Costs of bill of materials and supply chain currently unknown. Thanks to Medtronic for doing this! May not be suitable for ARDS: https://medium.com/@RobertLeeRead/medtronic-open-source-ventilator-does-not-meet-uks-rapidly-manufactured-ventilator-systems-81947c72a7ac |  | Some kind of PIC | Time-limited to Pandemic
Ambovent | https://1nn0v8ter.rocks/AmboVent-1690-108 | 4.5 | 4 | 4 | 4 | 3 | 4 | 4 | 3.9286 | 3 | 2020-04-01 | dreliram@gmail.com |  | AmbuBag |  |  | New | https://tinyurl.com/t2rp2y8 : Made open source on April 2nd. Extremely impressive documentation shown privately. Github org (no repos yet): https://github.com/AmboVent/AmboVent. No license by normal by US legal approach, but clearly dedicated to public domain in verbiage. Thanks to the AmboVent team and Isreal! | Hebrew, English, others | Arduino Nano | No known license, declared open source
MUR (Minimal Universal Respirator) | https://www.mur-project.org/ | 4 | 4 | 4 | 3 | 4 | 3 | 3.5 | 3.6429 | 2.5 | 2020-04-08 |  |  | Pressure Regulation |  |  | New | Bread boarding phase, look for documents (README minimal), some presure/volume testing. Not sure how the alarms work, but they have pressure sensor control. Exceptionally easy to use physical controls. Requires pressure sources. Valved solution produces some jitter in pressure wave; I don't know if that is clinically a problem or not. | French, English | Teensy 3.2
Rice OEDK Design: ApolloBVM | http://oedk.rice.edu/apollobvm/ | 5 | 4 | 4 | 2.5 | 2.5 | 3 | 2.5 | 3.3571 | 2.5 | 2020-04-03 | amy.k@rice.edu |  | AmbuBag |  |  | New | Note: Requires registration, but apparently open. Designs are CC-BY, code is GPL in repo: https://github.com/apollobvm/apollobvm 2020 effort informed by older effert, POC is E. D. of OEDK. Open sourced 04/03. Mechanical files are not in repo but downloadable in a zip file once you register. |  | Two Arduino Unos | CC-BY + GPL
COVID-19 Rapid Manufacture Ventilator (BVM) | https://www.instructables.com/id/COVID-19-Rapid-Manufacture-Ventilator-BVM-Ambubag-/ | 5 | 4 | 4 | 3 | 0 | 3.5 | 3.5 | 3.2857 | 2.5 | 2020-04-07 | https://www.instructables.com/id/COVID-19-Rapid-Manufacture-Ventilator-BVM-Ambubag-/ |  | AmbuBag |  |  | New | Part production can be fully automated via laser cutting, has a simple design, fast assembly, fully documented, but claims that it does not meet NHS standards
OpenVentilator (PopSolutions) | https://www.popsolutions.co/en_US/openventilator | 5 | 3.5 | 4 | 3.5 | 3 | 2 | 2 | 3.2857 | 3 | 2020-04-08 | contact@openventilator.io |  | Bellows |  |  | New | Great, well documented project; current test status hard to ascertian, but plans very good. Seems to use very widely available parts. | English, Portuguese
Open Source Ventilator Project | https://simulation.health.ufl.edu/technology-development/open-source-ventilator-project/ | 3.5 | 2.5 | 4 | 3.5 | 2 | 4 | 3 | 3.2143 | 4 | 2020-04-08 | https://simulation.health.ufl.edu/technology-development/open-source-ventilator-project/contact-us/ |  | Bellows |  |  | New | Best evidence of reliability. Seems very open, but has no license. Requires close watching!
A.R.M.E.E. Ventilator | www.armeevent.com | 3 | 5 | 4 | 2 | 3 | 2 | 2.5 | 3.0714 | 5 | 2020-04-06 | warrenkoch@gmail.com |  | Pnuematic |  |  | Existing | Also: https://docs.google.com/document/d/1por1ouQ4hEpHWrM--5hgncuU5_h0UzBLS1P5lqg7H7Q/edit?usp=sharing and has a gofundme: https://www.gofundme.com/f/highly-manufacturable-ventilators and repo: https://github.com/dogcomplex/oscillating_ventilator Large, active community at HelpfulEngineering | English | None
PREVAIL NY | https://jmawireless.com/prevail-ny/ | 4.5 | 4 | 3 | 2.5 | 0 | 4 | 3.5 | 3.0714 | 3.5 | 2020-04-09 | customerservice@jmawireless.com |  | AmbuBag |  |  | New | Open-source ambubag-type ventilator designed in collaboration with an MD. Provides a datasheet with full specifications, BOM, and instructions on assembly.
DIY-Beatmungsgerät [Respirator] | https://devpost.com/software/diy-beatmungsgerat | 5 | 4 | 3 | 1 | 2 | 3 | 3 | 3 | 0 | 2020-04-10 | https://docs.google.com/forms/d/e/1FAIpQLScmOmF4pyQQUXM5NEUqx-cSPIST0ftLQsrXMesnWElfhT6UqA/viewform |  | AmbuBag |  |  | New | In German, may be misreading. May not have PEEP, but mentions good alarm monitoring, so 3.5 on functional tests.
Low-Cost Open Source | https://github.com/jcl5m1/ventilator | 5 | 4 | 4 | 3 | 0 | 2 | 2 | 2.8571 | 3 | 2020-04-08 | https://github.com/jcl5m1/ventilator/issues, raise an issue |  | Pump |  |  | New | Has a good repo and community behind it with active projects
VentilAid | https://www.ventilaid.org/ | 5 | 4 | 4 | 3 | 0 | 2 | 2 | 2.8571 | 2.5 | 2020-04-08 | media@urbicum.com |  | AmbuBag |  |  | New | Great open source project that has amazing documentation that could be built to completion
Open Breath Italy | https://www.openbreath.it/ | 4 | 2.5 | 4 | 3 | 1 | 3 | 2 | 2.7857 |  | 2020-04-04 | info@openbreath.it |  | Pressure regulator |  |  | New | Hardware Repo: https://www.ohwr.org/project/openbreath/wikis/home Nice videos but hard to understand entires status. Not obvious what reliability or pressure/volume limit testing there is. Not obvious how the controller specifies rates. Lots of videos but it does not appear to me it would be possible to build from this repo and videos with out significant guesswork. Does seem to have BOM | Italian |  | CERN-OHL-S License
CoRescue | https://corescue.org | 4 | 1 | 4 | 3.5 | 0 | 3 | 3.5 | 2.7143 |  | 2020-04-15 | aau-pandemic-ventilator@aau.dk
MARK-19 Ventilator | https://www.mark-19.com | 3 | 2 | 3 | 4 | 3 | 2 | 2 | 2.7143 |  | 2020-04-15
Rice OEDK Design: (2019) ApolloBVM | https://docs.google.com/document/d/1-DRXnVkJOlDCmvTzh-DgWDxeLSrZTiBYyH0ypzv8tNA/edit | 3 | 3 | 3 | 3 | 2 | 3 | 2 | 2.71 |  | 2020-03-21 | amy.k@rice.edu |  | AmbuBag |  |  | New | Good documented solution from 2019 created for sneior design in university.
VentCore Ventilator | https://www.ventcore.health/ | 5 | 4 | 3.5 | 2 | 0 | 2 | 2 | 2.6429 | 3 | 2020-04-08 | https://www.ventcore.health/ |  | AmbuBag |  |  | New | Provides assembly files, production files, electronics, and schematics. Will be testing tidal volume and I:E rate with sensors in the near future.
Jeff Ebin's Prototpye | https://www.ebcore.io/?fbclid=IwAR3_SOqqJsohJ4rmR3tQVYS26_RNqx6A1nakRvklNlPNpJ4KS44iOL_a0RU | 5 | 4 | 3 | 1 | 0 | 2 | 3 | 2.5714 | 1 | 2020-04-08 | jeffrey@ebcore.io |  | AmbuBag |  |  | New | Created by MD, Worth a look
BLU3 VENT | https://www.diveblu3.com/blu3-vent/ | 1 | 2.5 | 2.5 | 3 | 3 | 3 | 3 | 2.5714 |  | 2020-04-10 | info@diveblu3.com |  | Pump |  |  | Existing (repurposed)
CoroVent | https://www.corovent.cz/ | 2 | 2 | 4 | 4 | 0 | 3 | 3 | 2.5714 |  | 2020-04-10 | simon.rakosnik@rubikonpr.cz |  | Bellows |  |  |  | They state that the device is on the prototype phase, and that they are testing components on lung models. They claim to publish the complete documentation, but thus far they don't appear to be open
A low oxygen consumption pneumatic ventilator for emergency construction during a respiratory failure pandemic | https://onlinelibrary.wiley.com/doi/full/10.1111/j.1365-2044.2009.06207.x | 4 | 3 | 2 | 4 | 1.5 | 1 | 2 | 2.5 |  | 2020-03-21 | j.dingley@swansea.ac.uk |  | AmbuBag |  |  | New | Very good research topic with layed out plans for pressure controls
INSPIRE - OpenLung | https://github.com/emersonmoretto/CITI-OpenLung | 5 | 2 | 4 | 3.5 | 0 | 3 | 0 | 2.5 | 3 | 2020-03-21 | https://docs.google.com/forms/d/e/1FAIpQLSfJmSH-XnRPcUmkUvS9ZdUzZxjUVUDR-olO96t6SzpphNBRpQ/viewform |  | AmbuBag |  |  | New | Developed a prototype that has undergone testign at a local university. They seem to provide updates daily, and are working on instructions for assembly. In Portuguese.
Worldwide Ventilator | https://www.worldwideventilator.com | 4 | 2.5 | 3 | 1.5 | 3 | 1.5 | 2 | 2.5 | 4 | 2020-04-04 | Kurt@Freshconsulting.com |  | Pneumatic |  |  | Existing | Beautiful reinvigoration of old Army projects, no moving parts. Printable parts not quite available fully but much information. Valuble addition to world-wide options due to simplicity and ease of manufacture. | English | None | Creative Commons Attribution-NonCommercial 4.0 International License.
Vortran G02VENT | https://www.vortran.com/go2vent | 1 | 2 | 2.5 | 4 | 0 | 4 | 4 | 2.5 |  | 2020-04-08 | info@vortran.com |  | Pneumatic |  |  | Existing | The device appears to be trademarked and not open source. Other groups have based their design off of this device. States that it does not require electrical power. They provide an excellent training video at youtube.com/watch?v=afFWwEclBW4
DRM127 Ventilator/Respirator | https://www.nxp.com/docs/en/application-note/DRM127.pdf | 1 | 2.5 | 1 | 5 | 0 | 4 | 4 | 2.5 |  | 2020-04-08 | nxp.com/support/support:SUPPORTHOME |  | Pump |  |  | Existing | Appears to be a 2011 ventilator model with specifications. Potentially not licensed.
Protofy Team OxyGEN | https://oxygen.protofy.xyz/ | 4 | 4 | 4 | 1 | 0 | 2 | 1 | 2.2857 |  | 2020-03-21 | https://www.oxygen.protofy.xyz/contact |  | AmbuBag |  |  | New | Easy mechanical build
S-VENT | https://s-vent.life/ | 4 | 2 | 2.5 | 2.5 | 0 | 3 | 2 | 2.2857 |  | 2020-04-07 | https://www.linkedin.com/in/tom-burton-1276877b/ |  | AmbuBag |  |  | New | Has videos of the working device, but no schematics. Made of wood.
crowdsourced-ventilator-covid-19 | https://github.com/Crowdsourced-Ventilator-Covid-19/crowdsourced-ventilator-covid-19 | 3 | 2 | 3 | 3 | 0 | 2 | 2.5 | 2.2143 |  | 2020-04-08 | chuck.tung@gmail.com |  | AmbuBag |  |  | New | Provides an instructional video on the function as well as design charg and software flowchart
LIBV | https://bvmvent.org | 4 | 2 | 3 | 0 | 0 | 4 | 2.5 | 2.2143 |  | 2020-04-17 | allendtan@gmail.com
respiArt | https://github.com/artur-iluk/respiArt | 4 | 1.5 | 2.5 | 2.25 | 0 | 2.5 | 2.5 | 2.1786 |  | 2020-04-12 |  |  | AmbuBag |  |  | New |  |  | Arduino
Illinois RapidVent | https://rapidvent.grainger.illinois.edu/ | 1 | 2.5 | 3 | 2.5 | 2.5 | 2 | 1.5 | 2.1429 |  | 2020-04-15 | Recently added, does not appear to be open, non-Ambubag design
Electric Blower Based Portable Emergency Ventilator | https://digitalcommons.usu.edu/cgi/viewcontent.cgi?referer=https://www.google.com/&httpsredir=1&article=1016&context=spacegrant&fbclid=IwAR1EtJVcxXm82PjGWFCA0t7H_MxNVjuseAePRfxNORr9h4ZQLQ9sNdQjXhc | 3 | 3 | 2 | 3 | 1 | 2 | 1 | 2.1429 |  | 2020-03-21 |  |  | Fan |  |  | New
The Open Ventilator | en.theopenventilator.com | 3 | 2 | 3 | 2 | 0 | 2.5 | 2 | 2.0714 |  | 2020-03-30 | info@theopenventilator.com |  | AmbuBag |  |  | New | Provides a video of the functioning prototype on a lung, along with controls.
BlueVent3d | facebook.com/BlueVent3d | 3 | 3 | 4 | 3 | 0 | 0 | 1.5 | 2.0714 |  | 2020-04-02 | facebook.com/BlueVent3d |  | Pump |  |  | New | They have several videos showing the functioning prototype, provide design files and a BOM.
OpenVent-Bristol V1.0 | https://www.instructables.com/id/COVID-19-Rapid-Manufacture-Ventilator-BVM-Ambubag-/ | 3 | 4 | 2.5 | 1 | 1 | 1 | 2 | 2.0714 | 4 | 2020-04-06 |  |  | AmbuBag |  |  | New | Very simple, very easy to make. Good instructions. Easily sourced parts. I see this is a good starting point for other projects. To improve it needs more testing.
Zephyr Open Source Ventilator | https://github.com/Julhh/Zephyr | 3 | 2 | 3 | 2 | 0 | 3 | 1.5 | 2.0714 |  | 2020-04-08 | juliencope@gmail.com | [5 empty] | Code is provided, as well as videos detailing the CAD design and a functional prototype |  | Arduino
The AIR Project | https://the-air-project.github.io | 3 | 2 | 3.5 | 0 | 0 | 4 | 2 | 2.0714 |  | 2020-04-10 |  |  | AmbuBag |  |  |  |  | English
MIT 2010 (Husseini et al) | https://web.mit.edu/2.75/projects/DMD_2010_Al_Husseini.pdf | 2 | 2 | 2 | 3 | 1 | 2 | 2 | 2 |  | 2020-03-28 |  |  | Pump
CaRE-VENT | https://care-vent.web.unc.edu/home/testing/ | 3 | 1 | 2 | 3 | 0 | 3 | 2 | 2 |  | 2020-04-07 | care-vent@unc.edu |  | AmbuBag |  |  |  | Has videos and testing documentation but does not provide instructions on assembly
Saving Babies' Lives Starts With Aquarium Pumps And Ingenuity | https://www.npr.org/sections/health-shots/2014/01/03/259436844/saving-babies-lives-starts-with-aquarium-pumps-and-ingenuity | 1 | 1 | 2 | 5 | 0 | 0 | 4 | 1.8571 |  | 2020-03-21 |  |  | Pump |  |  |  | CPAP for infants
Gtech Ventilator | https://www.gtech.co.uk/ventilators | 4 | 4 | 3 | 0 | 0 | 2 | 0 | 1.8571 | 2.5 | 2020-04-08 | ventilator@gtech.co.uk |  | AmbuBag |  |  |  | They provide design files and instructions as well as assembly and calibration videos. Not sure if entirely open.
MIT E-Vent | https://e-vent.mit.edu/ | 2.5 | 1 | 3.5 | 2 | 0 | 3 | 0 | 1.7143 |  | 2020-04-08 |  |  | AmbuBag |  |  |  | Apparently tested on a live porcine model |  |  | MIT Open Source license, found here: https://opensource.org/licenses/MIT.
VentilatorPAL | https://freebreathing.org/ | 2 | 1 | 3 | 1.5 | 1 | 1.5 | 2 | 1.7143 |  | 2020-03-21 | https://freebreathing.org/ |  | AmbuBag |  |  |  | Pre-orderable. Large but somewhat opaque team, little test information. |  | Arduino Nano
Open source ventilator Pakistan | https://docs.google.com/spreadsheets/d/1PfxblJfeVo9WGIUSiz_yQkhYF2OCvRRPal1zllWFMHY/edit#gid=1040995025 | 3 | 1 | 2 | 0 | 0 | 3 | 3 | 1.7143 |  | 2020-03-21 | https://docs.google.com/spreadsheets/d/1PfxblJfeVo9WGIUSiz_yQkhYF2OCvRRPal1zllWFMHY/edit#gid=1746079259 |  | AmbuBag |  |  |  | Well documented, large community, but in design phase
openventilator - KiCad Translation and update of the Medtronic OpenVentilator | https://gitlab.com/openventilator/openventilator | 4 | 4 | 4 | 0 | 0 | 0 | 0 | 1.7143 |  | 2020-04-09 | seth@kipro-pcb.com | [5 empty] | Provides updated and modernized KiCAD files for Medtronic's open ventilator and a BoM to parts that are available |  |  |  |  | Plan to seek emergency approval soon
Open Source Ventilator - OpenLung BVM Ventilator | https://gitlab.com/open-source-ventilator/OpenLung | 5 | 2.5 | 4 | 0 | 0 | 0 | 0 | 1.6429 |  | 2020-04-15 | https://opensourceventilator.ie/register |  | AmbuBag
CoronavirusMakers | https://gitlab.com/coronavirusmakers | 3 | 2 | 3 | 0 | 0 | 2 | 1 | 1.5714 |  | 2020-03-21 | https://www.coronavirusmakers.org/index.php/es/54-telegram-3 | [5 empty] | In Spanish, can only read a little, may have more
The Pandemic Ventilator (older) | https://www.instructables.com/id/The-Pandemic-Ventilator/ | 3 | 4 | 1 | 1 | 0 | 1 | 1 | 1.5714 | 1 | 2020-04-08 |  |  | Bellows
Cuirass-Ventilator | https://github.com/jps2000/Cuirass-Ventilator | 3 | 2 | 3 | 2 | 0 | 0 | 1 | 1.5714 |  | 2020-03-28 | https://github.com/jps2000/Cuirass-Ventilator/issues |  | Cuirass |  |  |  | Cuirass type ventilator, clinical suitability unevaluatable
SparkVent | https://github.com/SparkVent/TheSparkVent | 4 | 4 | 3 | 0 | 0 | 0 | 0 | 1.5714 |  | 2020-04-08 |  |  | Bellows
YACoVV - Yet Another (SARS-)CoV(-2)Ventilator | https://github.com/auenkind/YACoVV | 2 | 1 | 2.5 | 1 | 0 | 3 | 1 | 1.5 |  | 2020-03-21 |  |  | Water column |  |  |  | Fluid pressure based solution
IMPROV: Inexpensive Maker-Made Piston-Respiratory Open-Source Ventilator | https://www.izutadesign.com/ | 4 | 2 | 3 | 1.5 | 0 | 0 | 0 | 1.5 |  | 2020-03-30 | Design@izutadesign.com |  | Pump |  |  |  | Most of the documents they provide describe the design rather than specify it. They performed some functional tests on a piston system rather than a lung or lung simulation.
Ad Hoc Ventilator | https://20100.be/ad-hoc-ventilator/ | 2.5 | 2 | 2 | 1 | 0 | 2 | 1 | 1.5 |  | 2020-04-03 | https://20100.be/contact-2/ |  | AmbuBag |  |  |  | Provides a video of the working device. Essentially an automatic ambu-bag pump
Tahl Inc. Ventilator | https://www.youtube.com/watch?v=uacn6mHcFPk | 2 | 1 | 2.5 | 0 | 0 | 3 | 2 | 1.5 |  | 2020-04-14 | shop@tahlinc.com |  | AmbuBag
RespiraWorks | https://respira.works/ | 1 | 1 | 3 | 3 | 0 | 2 | 0 | 1.4286 |  | 2020-03-21 | ethan@respira.works |  | Fan |  |  |  | Good document for design, not finished, a work in progress
MIT Low Cost Ventilator | https://github.com/RuairiSpain/openVentilator | 3 | 2 | 2 | 1 | 0 | 0 | 2 | 1.4286 |  | 2020-03-21
Dr. Mujeeb ur Rahman design | http://www.technologyreview.pk/pakistani-engineer-braves-tragedy-to-develop-low-cost-ventilator/ | 1 | 1 | 2 | 1 | 0 | 2 | 3 | 1.4286 |  | 2020-03-21 |  |  | AmbuBag
Hackaday Rex Ventilator V1 | https://github.com/hackadayrex/ventilator | 2.5 | 2 | 2.5 | 1 | 0 | 1 | 1 | 1.4286 |  | 2020-03-21 | hackadayrex@gmail.com |  | Bellows |  |  |  | Great photos and video but no reproducible plans.
PulmoLogic | https://pulmologic.com/#clipvent | 1 | 1 | 2.5 | 2.5 | 0 | 3 | 0 | 1.4286 |  | 2020-04-16 | https://pulmologic.com/#support |  | Pneumatic |  |  |  | States that the device will be available pending final development stages and validation
Automatic ambu ventilator | https://github.com/SebaGaggino/rees31 | 3 | 2 | 2 | 0 | 0 | 1 | 1.5 | 1.3571 |  | 2020-03-29
Pandemic Ventilator | https://www.cbc.ca/news/canada/london/pandemic-ventilator-coronvirus-hospitals-1.5493830 | 1 | 1 | 2 | 2 | 0 | 2 | 1 | 1.2857 |  | 2020-03-21 |  |  | Pump |  |  |  | Inventor says may open source
Open Ventilator Project | https://docs.google.com/document/u/1/d/e/2PACX-1vSrCNygk8LHYIgvpu8tXlVVZrSniwSC7HSBAhMUoOsnFckagZwPCVyNStLS7wBVwR9K0gvYXtRynPKx/pub | 3 | 2 | 4 | 0 | 0 | 0 | 0 | 1.2857 |  | 2020-03-30 |  |  | Fan
OpenVentilator | https://github.com/OpenVentilator/OpenVentilator | 3 | 2 | 3 | 0 | 0 | 1 | 0 | 1.2857 |  | 2020-04-03 |  |  | AmbuBag |  |  |  | Not built yet, seems to be design phase
Simple device from www.POMO.cl | https://www.facebook.com/groups/670932227050506/permalink/675264606617268/?app=fbl | 1 | 2 | 3 | 1 | 0 | 0 | 1 | 1.1429 |  | 2020-03-21 | Mserey@pomo.cl |  | AmbuBag
Acute-19 | https://valenciaplaza.com/desarrollan-ventilador-turbina-impresion-3d | 2 | 1 | 2 | 0 | 3 | 0 | 0 | 1.1429 |  | 2020-03-30 | info@acute19.com | [5 empty] | Reports testing for 48 hours on various lung models. Does not specify parameters or provide details. Will be tested on sheep in the nearby future. Provided website rather empty acute19.com/blog/
COVID19 Respirador (Vaccarini) | https://github.com/vaccarini/covid19respirador | 3 | 2 | 3 | 0 | 0 | 0 | 0 | 1.1429 |  | 2020-03-30 | martin.vaccarini@gmail.com
The Breathing Project | https://www.uni-marburg.de/de/fb13/halbleiterphotonik/the-breathing-project/the-breathing-project-1 | 1.5 | 1 | 3 | 2.5 | 0 | 0 | 0 | 1.1429 |  | 2020-03-30 | breathing.project@physik.uni-marburg.de |  | AmbuBag/Fan |  |  |  | Provide videos of prototype testing. Controls are visible, but function not demonstrated. State an eventual goal of having the design be public, but have not published any open designs or instructions yet.
Cuirass Ventilator the DIY way | https://www.youtube.com/watch?v=pvrUQCMa3a8&feature=youtu.be | 2 | 2 | 1 | 1 | 0 | 0 | 1 | 1 |  | 2020-03-21 |  |  | Cuirass
1M Ventilators | https://1mventilators.com/ | 2 | 1 | 3 | 0 | 0 | 1 | 0 | 1 |  | 2020-03-21 | https://1mventilators.com/community/ |  | AmbuBag |  |  |  | Repo: https://github.com/1M-Ventilators, just beginning
MVP | https://drive.google.com/drive/folders/1XTNc5_pEWI1ATr-WRFbC4TW5NCqKutzC?usp=sharing | 4 | 1 | 0 | 0 | 0 | 1 | 1 | 1 |  | 2020-03-29 | yoram@scope5.com |  | AmbuBag
Open Source Ventilator Ireland | https://opensourceventilator.ie/ | 4 | 2 | 0 | 0 | 0 | 0 | 1 | 1 |  | 2020-03-28 | opensourceventilatorireland@gmail.com |  | AmbuBag
Breating Aid | https://www.breathing-aid.org/homeen | 2 | 2 | 2 | 0 | 0 | 0 | 1 | 1 |  | 2020-03-30 | covid@nemos.org |  | Fan |  |  |  | Seems to be CPAP for multiple patients; could be useful as relief for equipment shortage for non-ARDS patients?
Low-Cost-Medical-Ventilator | https://github.com/snipe75/low-cost-medical-ventilator | 2.5 | 2 | 1 | 1.5 | 0 | 0 | 0 | 1 |  | 2020-04-02 |  |  | Fan |  |  |  | Don't have build plans but provide their functional code. Works as an ambu bag-type ventilator
Venezuelan Ventilator | https://www.instagram.com/p/B-n3SiaHEOY/?igshid=laz6eje29rm2 | 2 | 1 | 4 | 0 | 0 | 0 | 0 | 1 |  | 2020-04-15 | pbcelis@usb.ve & ftorre@usb.ve | A lot. Specially electronic components and measuring devices | AmbuBag |  |  |  | We are working on a mechanical device to pressure the AmbuBag (MIT like) and another using pneumatic pressure from air/oxygen line. It is a device done with what we may find available, which is scarce. | Spanish
Pandemic Ventilator Project | https://panvent.blogspot.com/2008/02/test-of-pandemic-ventilator-with.html | 2 | 1 | 1 | 2 | 0 | 0 | 0 | 0.85714 |  | 2020-03-28 |  |  | Bellows
Mechanical Ventilator Milano (MVM) | https://arxiv.org/pdf/2003.10405.pdf | 1 | 2 | 1 | 1 | 0 | 0 | 1 | 0.85714 |  | 2020-03-21
OxVent | https://oxvent.org/ | 1 | 1 | 1 | 3 | 0 | 0 | 0 | 0.85714 |  | 2020-03-28 | https://oxvent.org/contact/ |  | AmbuBag |  |  |  | Could be great, but almost no information
Automatic Rescucitator | https://www.thingiverse.com/thing:4239988/files | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0.85714 |  | 2020-03-30 | https://www.frax3d.com/contactar |  | AmbuBag |  |  |  | Contact through form on https://www.frax3d.com/contactar. No specifications, mostly just provides 3D models of parts.
Open Source Covid-19 Ventilator Canada | https://open-source-covid-19-ventilator-canada.mn.co/ | 2 | 1 | 3 | 0 | 0 | 0 | 0 | 0.85714 |  | 2020-03-30
Vortran-Type Pneumatic Ventilator | https://docs.google.com/document/d/1R25O2mKT3TfSjXTmheGEevIk6rTJ49o-sGCJU3QqP3M/edit# | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0.85714 |  | 2020-03-30 | [6 empty] | Based off of a commercial design
Low Cost Medical Ventilator | https://github.com/snipe75/low-cost-medical-ventilator | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 0.85714 | 0 | 2020-04-06 |  |  | Fan |  |  |  | Very early project
Project DIY Ventilator | https://github.com/Mayoogh/Project-DIY-Ventilator | 3 | 2 | 1 | 0 | 0 | 0 | 0 | 0.85714 |  | 2020-04-16
Low-Cost Automated Emergency Ventilator | https://gitlab.com/PerAsperaAdAstra/lcaev | 1 | 1.5 | 2 | 1 | 0 | 0 | 0 | 0.78571 |  | 2020-03-30 |  |  | AmbuBag |  |  |  | Only understandable by view design log: https://gitlab.com/PerAsperaAdAstra/lcaev/-/wikis/Design%20Log
Low-Cost Ventilator Wins Sloan Health Care Prize | https://www.medicaldesignandoutsourcing.com/low-cost-ventilator-wins-sloan-health-care-prize/ | 1 | 1 | 1 | 0 | 0 | 2 | 0 | 0.71429 |  | 2020-03-21 |  |  | AmbuBag |  |  |  | Very little information
Projecto EAR Celso | https://github.com/RespiradorHacker/Projeto-EAR-Celso | 2 | 1 | 0 | 0 | 0 | 0 | 2 | 0.71429 |  | 2020-03-21 | https://github.com/RespiradorHacker/Projeto-EAR-Celso |  | AmbuBag
Project Open Air | https://www.projectopenair.org/ | 1 | 1 | 0 | 0 | 0 | 0 | 3 | 0.71429 |  | 2020-03-28 | openair@lift.com.pt
LEITAT1 Respirator | https://www.3dnatives.com/en/3d-printed-respirator-230320205/ | 1 | 2 | 2 | 0 | 0 | 0 | 0 | 0.71429 |  | 2020-03-30 | lluis.lopez@zfbarcelona.es |  | AmbuBag
Respirador RESP19 | https://www.gofundme.com/f/respirador-resp19?utm_medium=chat&utm_source=whatsapp-visit&utm_campaign=p_lico+share-sheet | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 0.71429 |  | 2020-03-30 | https://www.gofundme.com/f/respirador-resp19?utm_medium=chat&utm_source=whatsapp-visit&utm_campaign=p_lico+share-sheet |  | Fan |  |  |  | Thus far it's just an idea
OperationAIR | https://www.tudelft.nl/onderwijs/opleidingen/masters/technical-medicine/msc-technical-medicine/operationair/ | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 0.71429 |  | 2020-03-30 | info@operationair.org | [5 empty] | State that the design and information will be published open source
Inspiramed | https://letsbeatcovid19.blogspot.com/2020/03/march-29th.html | 2 | 2 | 0 | 1 | 0 | 0 | 0 | 0.71429 |  | 2020-04-04 |  |  | Pressure? |  |  |  | Seems to be well-tested, but seems to have only video (though with schematics.) | Spanish
Ventilador Foscal y Unab | https://www.vanguardia.com/area-metropolitana/bucaramanga/esta-listo-el-primer-ventilador-mecanico-fabricado-100-en-santander-BG2181982 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0.57143 |  | 2020-03-30 |  |  | AmbuBag |  |  |  | Very little information, provides a video showing the device pumping
The Coventor | https://med.umn.edu/covid19Ventilator | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0.42857 |  | 2020-04-17 |  |  | AmbuBag | [8 empty] | FDA approved
License: | CC0
Copyright Robert L. Read, Keeshan Patel, Juan E. Villacres Perez and Avinash Baskaran, 2020
Rank by Average

## Drives
Drive (Air/Oxygen pressure and flow provision)
Drive Module (name) | Project Link | Openness | Community Support | Ease of Integration | Buildability (1 unit) | Manufacturability (1000s) | Average | Unit Costs | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Notes
BLU3 Vent | https://www.diveblu3.com/blu3-vent/ | 1 | 2 | 1 | 2.5 | 4 | 2.1 | $1000? | 2020-04-11 |  | None? | Not open source, but approachable. They have a diagham pump designed for and excellent for this purpose. Their pump can be essentially treated as an independent controllable module.

## Monitors
Monitoring Modules (includes alarms and sensing modules w/ and w/out UIs)
Drive Module (name) | Project Link | Openness | Community Support | Ease of Integration | Buildability (1 unit) | Manufacturability (1000s) | Average | Unit Costs | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Licensing | Notes | Additional Notes
VentMon | https://www.diveblu3.com/blu3-vent/ | 5 | 4 | 4 | 3 | 2 | 3.6 | 250 | 2020-04-14 | read.robert@gmail.com | Crack programmers | MIT, GPL, and CC0
FreeFlow flow sensor | https://hackaday.io/project/170928-freeflow-flow-sensor | 5 | 4 | 4 | 3 | 2 | 3.6 | 50 | 2020-04-15 |  |  | CERN open hardware | Awesome project, includes 3D printable venturi tube. Will try to integrate with VentMon. | https://github.com/sglow/freeflow

## Flow Sensors
Flow Sensors (as both components and modules, may include pressure sensors)
Drive Module (name) | Project Link | Openness | Community Support | Ease of Integration | Buildability (1 unit) | Manufacturability (1000s) | Average | Unit Costs | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Licensing | Notes | Additional Notes
VISP -- Ventilator Inline Sensor Package | https://hackaday.io/project/170622-visp-ventilator-inline-sensor-package | 5 | 4 | 4 | 3 | 2 | 3.6 | 50 | 2020-04-14
Sensirion sensors |  | 2 | 3 | 5 | 0 | 0 | 2 | ~$200 | 2020-04-14 |  |  |  | A commercial product by Sensirion, a good corporate citizen. Out-of-stock world-wide, ramping to 6x annual production in 2 quarters. No direct sales less than groups of 1000
FreeFlow flow sensor | https://hackaday.io/project/170928-freeflow-flow-sensor | 5 | 4 | 4 | 3 | 2 | 3.6 | 50 | 2020-04-15 |  |  | CERN open hardware | Awesome project, includes 3D printable venturi tube. Will try to integrate with VentMon. | https://github.com/sglow/freeflow | https://github.com/sglow/freeflow
DIY Beatmungsgerät Mass Air Flow Meter (also supporrting ready-made Sensirion sensors, compatible digitial interface (I2C)), reusable/modular | https://github.com/maehw/diy-beatmungsgeraet/blob/master/software/mass-airflow-sensor/documentation/README.md | 5 | 4 | 4 | 4 | 4 | 4.2 | ~$50 | 2020-04-18 |  |  | GPL | This needs more study but appears to be a completely servicable solution, and completely reproducible. I am convinced this approach in principle solves the problem.

## Standards
Standards (Generally Vent. specific, which support interoperability, cooperation, and modularity.)
Note: These may be software, electrical (e.g. I2C), or some combination
Drive Module (name) | Project Link | Openness | Community Support/Testability | Ease of Integration | Average | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Notes
PIRDS | https://github.com/PubInv/respiration-data-standard | 5 | 4 | 4 | 4.3333 | 2020-04-14 | read.robert@gmail.com |  | A standard for recording specific sensor measurements, most obviously for breath traces

## User DisplayControl
User Display/Control Panels (physical controls, may overlap software)
Note: This tab was created on April 14th and is in an early stage or readiness.
Drive Module (name) | Project Link | Openness | Community Support | Ease of Integration | Buildability (1 unit) | Manufacturability (1000s) | Average | Unit Costs | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Licenses | Notes
Delphi | https://github.com/jennyfilipetti/delphi-emergency-ventilator-controller | 5 | 3 | 2 | 1 | 2 | 2.6 | $50? | 2020-04-15 | jkb@colorado.edu |  | GPL
VentSim | https://github.com/CSSALTlab/Open_Source_Ventilator/tree/master/ControllerSourceCode/VentSim | 4 | 3 | 2 | 1 | 2 | 2.4 | $50? | 2020-04-15 |  |  | None stated | Evaluation early. CSSALTlab project. Seems open but no license file. Hard to tell progress from the repo.This appears to be separation of controller from the UI, a great idea.
ArduinoVent | https://github.com/CSSALTlab/Open_Source_Ventilator/tree/master/ControllerSourceCode/ArduinoVent | 4 | 3 | 2 | 1 | 2 | 2.4 | 40? | 2020-04-15 |  |  | None stated | Evaluation early.CSSALTlab project. Seems open but no license file. Hard to tell progress from the repo.This appears to be separation of controller from the UI, a great idea.

## Oxygen Blender
Oyxgen Mixer/Blenders (tools for mixing and reducing pressures of medical gases, including sensors)
Drive Module (name) | Project Link | Openness | Community Support | Ease of Integration | Buildability (1 unit) | Manufacturability (1000s) | Average | Unit Costs | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Notes
Fio2mix | https://github.com/aizvorski/fio2mix | 5 | 4 | 4 | 3 | 2 | 3.6 | $10? | 2020-04-14 | Alex Izvorski |  | This is an entrainment nozzle for adding high-pressure oxygen gently to a flow of air.

## Sanitizer
Sanitizers (for the end of the exhalation circuit, specifically to reduce viral load of exhalation)
Drive Module (name) | Project Link | Openness | Community Support | Ease of Integration | Buildability (1 unit) | Safety for Clinicians | Ease of clinical Use | Manufacturability (1000s) | Average | Unit Costs | Date Last Evaluated | Point of Contact | Needs (needs of the teams and involvement opportunities) | Notes
no yet know |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Mentionable
Useful Projects Which are Not Ventilators
Michigan Tech Open Sustainability Lab has a Test Lung that can do sophisticated testing if you ship them a ventilator:
https://www.appropedia.org/Artificial_lung_method_MOST
https://github.com/PubInv/ventmon-ventilator-inline-test-monitor
VentMon is Public Invention's attempt to build an inline tester and Critical Care Monitor for any Ventilator
This appears to be an oxygen mixer, not a ventilator, but someont should study:
https://www.mercedesamgf1.com/en/news/2020/04/ucl-uclh-f1-project-pitlane-start-delivery-breathing-aids-nhs-hospitals/
VISP - Ventilaotr Inline Sensor Package
https://hackaday.io/project/170622-visp-ventilator-inline-sensor-package
This is a great project: they are 3D printing an aiway tube-compatible flow and pressure sensor in a single pacakge.
The following projects are mentionable, but do not have enough informaton to evalue:
From March 27. 2020
Omar from Colombia | https://diyventilators.discourse.group/t/first-prototypes-by-omar-from-lif-and-team/112/2
"has built a functional prototype of the Charlotte Valve, developed in Italy, but modified to be built with scuba gear."
Bozhana's group from Resonator.bg has build the Decathlon mask, with a modification for using it as PPE, and it should be tested in a Bulgarian hospital today or tomorrow. She needs visibility on the level of stock of CPAP or sleep apnea machines in Bulgaria to assess her next steps. | https://diyventilators.discourse.group/t/decathlon-mask-for-cpap-and-for-medical-staff-protection/75
Bozhana's group from Resonator.bg has build the Decathlon mask, with a modification for using it as PPE, and it should be tested in a Bulgarian hospital today or tomorrow. She needs visibility on the level of stock of CPAP or sleep apnea machines in Bulgaria to assess her next steps.
Raod's team
Rado's team has built a prototype using a combination of the MIT e-vent and the 3D printed respirator that was approved in Spain. They're working on a PCB controller (not arduino) to make certification easier. They need immediate help from people experienced with certification and validation.